| ## Name | ## CALEX |
| --- | --- |
| Author | Erhard Wielandt, Institute of Geophysics, University of Stuttgart, Richard-Wagner-Strasse 44, D - 70184 Stuttgart, E-mail: e.wielandt@t-online.de |
| Version | 2007 |

This program determines system parameters, such as corner periods, damping constants, poles and zeros of the transfer function of analog filters or seismometers, by modeling the response of the system to an arbitrary input signal in the time domain, and fitting it to the observed output signal. The input signal and the response must be digitally recorded. The test signal may be generated by any practical means, even by operating a switch or a potentiometer by hand; it must however contain sufficient energy at frequencies of interest.

The program requires three input files in ASCII format:

-	A file ‘calex.par’ containing numerical parameters for the inversion and start values for

the system parameters. A different name for the parameter file can be entered as

a runstring.

-   An input file, here referred to as ‘eing’, containing the test (input) signal

-   An output file, here referred to as ‘ausg’, containing the output signal

The names of the input and output files are specified in the calex.par file.

The following output files are generated:

’einf ’ :	a low-pass (anti-alias) filtered version of ‘eing'

‘ausf ' :	a low-pass filtered version of ‘ausg’ for direct comparison with ‘synt’

‘synt ’:	the synthetic output signal. If this does not fit the signal

‘ausf ' at all you have probably made a mistake in defining the system.

’rest ’:	the residual. These files have the same format as ‘eing’ and ‘ausg’

‘calex.out ’: a protocol of the inversion, identical to the printout on the screen. Its central section lists the number of the iteration, the rms residual error (normalized to the rms output signal), and the normalized values of the active parameters. The absolute value of each parameter (as given at the end of the inversion) is:

start value + normalized value * uncertainty.

Format of the DATA FILES

The structure of the data files is as follows:

- one header line, arbitrary (will be echoed but not evaluated)

- one line containing the number of samples, the FORTRAN format in which they are listed, and the sampling interval. These three entries must be in the FORTRAN format (i10,a20,f10.x).

- data in the specified format

You may easily change this file structure in the source code.

Format of the PARAMETER FILE

Calex.par contains six types of lines:

- a header line

- two lines with the names of the input (stimulus) and output files. File names should be enclosed in single quotes beginning in the first column.

- ten lines with a control parameter (a single number) each for the inversion

- four lines specifying required parameters for the inversion

- optional lines with a keyword defining a subsystem

- optional lines specifying parameters for the subsystem (one line for each parameter)

- the end line

These lines appear in the sequence: header – file names - control parameters – required parameters - keyword - subsystem parameter(s) - keyword - subsystem parameter(s) - … -  end

Lines beginning with a blank are considered as comments and ignored. The control parameters (one parameter per line, flush left) are:

alias:	the corner period of the numerical anti-alias filter that is part of the CALEX routine. Must be at least 4-5 times larger than the sampling interval. This filter is required because the bandwidth of the simulated system must be smaller than the Nyquist bandwidth. The program determines the order (steepness) of the filter so that the Nyquist condition is satisfied.

m:	the number of active (unknown) parameters in the inversion. The set of parameters comprises a gain factor, a constant delay, two parameters for the correction of  experimental biases, plus the unknown corner periods and damping constants. Any selection of these may be declared active, that is, be specified with a nonzero uncertainty. The program will fit only the active parameters and leave the passive unchanged.

m0:	number of additional powers of the Laplace variable s in the nominator of the transfer function, equivalent to forming the m0-th time derivative of the signal. m0 may also be < 0 when the signal is integrated in the system.

ml:      number of first-order high-pass or low-pass subsystems in the transfer function. A corner period must subsequently be specified for each of these.

m2:      number of second-order high-pass, low-pass or band-pass subsystems in the transfer function. A corner period and a damping constant (fraction of critical damping) must subsequently be specified for each of these.

maxit:    the maximum number of iterations in the conjugate-gradient optimization procedure

qac,

finac:		the iteration stops when the improvement in the rms misfit in one step becomes less than qac and the normalized parameters change by less than finac.

nsl,

ns2:	the input and output signals are analyzed in a time window from sample nsl to sample ns2. The values 0 mean first resp.last sample. If the system was not quiet at the beginning of the record, ns1 must be chosen so that the transient response to all previous input is excluded from the analysis.

The system parameters – both required and optional - are specified as follows: a line with

nam val unc (three entries separated by spaces) where

nam is an arbitrary name (three characters) that appears in the printout

val is the initial value of an active parameter, resp. the fixed value of a passive parameter

unc is the estimated uncertainty of an active parameter resp. zero for a passive parameter

Four system parameters are required:

amp	a gain factor (not the generator constant of the sensor). The start value may be

estimated from the amplitude ratio between output and input.

The meaning of the 'amp' parameter is different for different experiments

(electrical or mechanical input, etc.). If a broadband-velocity seismometer

with the asymptotic (practically, mid-band) generator constant G is tested

on a shake table that has a displacement transducer, two possible cases are:

a) the displacement transducer has an electric output of D volts per meter

and its output signal is recorded with the same digitizer as the seismo-

meter output. Then amp = G / D.

b) The displacement transducer is digital, or has a separate digitizer, with

a responsivity A counts per meter. The seismometer output is digitized

with C counts per volt. Then amp = C * G / A.

del	a time delay, normally used to describe the delay of low-pass filters in the system that

are not explicitly specified. Any skew in the sampling of the digitizer channels, or other differential delays, may also be included here. If you don’t want to use this parameter, set its initial value and uncertainty to zero. Its use is however recommended because this improves the mathematical model of your system.

sub	the fraction ‘sub’ of the input is subtracted from the output, to numerically simulate a

full bridge if a half-bridge circuit was used. Set start value and uncertainty to zero if  no bridge circuit was used.

til	This parameter permits to subtract ‘til’ times the twice-integrated synthetic output

from the synthetic output signal, to compensate for the tilt of a shake table. Set

start value and uncertainty to zero if you did not use a shake table.

til is in units of microradians per millimeter of table motion provided

that the motion of the table, the gravitational force coupled in by tilt,

and the sensitive axis of the seismometer have the same direction. More

generally, til is (tilt along the sensitive axis) per (table displacement

along the sensitive axis), in the above units.

The number and arrangement of the optional system parameters depends on the system. Specify all first-order subsystems before all second-order subsystems. The parameter lines for each subsystem must be preceded by a line defining the type of the subsystem: lp1, hp1, 1p2, bp2, or hp2 where lp1 denotes a low-pass filter of first order, etc. Example:

lp2

per  30.       1.

dmp	0.707     0.01

for a second-order low-pass subsystem. The damping parameter appears only in second-order subsystems. The names must not be blank. Note that the parameter name has three characters, followed by two spaces. Start parameters and their uncertainties occupy ten spaces each.

You may include any (reasonable) number of additional subsystems such as analog filters in the analysis, by simply appending their definition and their parameters to the calex.par file. The parameters m, ml and m2 must be specified accordingly. The program will check the number of active parameters (defined by nonzero uncertainty) and the number of subsystems of first and second order, and complain if there is a discrepancy.

In the unlikely event that you want to use more than 24 system parameters, more than 12 active parameters, or more than 480 000 data samples, you must change the array dimensions in the source code and recompile it.

Example for a parameter file

A complete parameter file is given below. The seismometer is described as a bandpass filter (which it actually is when excited over the calibration coil). Alternatively it could be described as a low-pass filter with an additional differentiation (m0=1), or as a high-pass filter with an additional integration (m0= -1). The gain factor amp would be different in these three cases. Note that the gain is only a fitting parameter here, it is not the generator constant of the sensor.

Calibration of the broadband (20 sec) seismometer STS-1 No.14

‘eing’   input file

‘ausg’   output file

1	alias

4	m

0	m0

0	ml

1	m2

48	maxit

le-5	qac

le-2	finac

1	ns1

0	ns2

amp  1.3       0.10

del  0.01      0.01

sub	0.        0.

til	0.        0.

bp2

per  20.       1.

dmp  0.7       0.1

end

##### Download

Download source code, executable, and test data from

http://www.software-for-seismometry.de/