| ## Name | ## SINCAL1, SINCAL2 |
| --- | --- |
| Author | Erhard Wielandt, Institute of Geophysics, University of Stuttgart, Richard-Wagner-Strasse 44, D - 70184 Stuttgart, E-mail: e.wielandt@t-online.de |
| Version | 2005 |

These programs determine the three parameters of an electromagnetic seismometer: free period, damping and generator constant (gain) from a record of 4 to 9 test frequencies. It cannot be used for systems with a more complicated response. SINCAL1 requires the test frequencies to be precisely known while SINCAL2 determines them from the recorded output signal, so they must not be precisely known, and an analog sine-wave generator may be used.

The programs assume that you have recorded the response of a

simple seismometer (with a geophone-type transfer function) to

at least four sinusoidal test signals of equal amplitude. When

f0 is the eigenfrequency of the sensor, then the signal frequencies

0.2*f0, 0.5*f0, f0, 2*f0, and 5*f0 are a good choice. The amplitudes

must be equal in terms of calibration coil current or ground

acceleration. If your calibration-signal generator puts out

signals corresponding ground velocity, the amplitudes must be

chosen inversely proportional to the frequency. Approximate values

of the test frequencies must be communicated to the program as the

first parameter t0 of the 'ban' commands in the 'sincal5.par' file.

Exact values are then determined automatically.

The records of the test signals must be contained in a single

data file. Each test signal should contain about 25 full periods,

depending on the half-width of the gaussian bandpass filters used

in the analysis. This half-width, measured in periods, is entered

as the second parameter, m, of the 'ban' command in the sincal5.par

file. m=3 is generally sufficient; higher values result in more

selective filters, thus a better immunity to noise. Each test

frequency must be present for at least 5*m periods at constant

amplitude (more if tapered). Between the test signals, quiet

intervals corresponding to m periods of the slower one of the

adjacent signals should be present. These requirements  may be

relaxed – just try.

The program is controlled by a parameter file 'sincal5.par'

unless a different file name is specified in the runstring.

It has limited time-domain  signal processing capabilities

like SEIFE, but is not a general data-processing tool, and

produces no output signal.

Convention for the 'seife' type file format:

- one header line, up to 72 characters

- up to 48 lines of signal description and processing log

these lines have '%' as the first character

- one line specifying the number of samples, the format, the sampling

- data in any FORTRAN format

Other supported formats are:

- single-column ASCII with header lines ignored (read only)

- MARS88 - ASCII with header line (read only)

Consult the SEIFE program for details (comments are part of the source code

and can be printed with ‘seife info’)

The desired time-domain processes are defined in a command file

whose name is entered either interactively or as a runstring.

this file also specifies the names of the input files (up to 24; all

processes apply to all input files). each line of the command file

consists of a thee-letter code followed by one or more blanks and some

numerical parameters (such as corner period, damping, window limits).

parameters must be separated by blanks or commas and fit into columns

5 to 50.

the following codes are accepted:

blank first character, or rem: line is ignored

commands for input and output:

fil  nam1: specify input file names. Will read 'seife' type files.

mar  nam1  like fil, but for MARS88-ASCII files from MARSDUMP

asc  nam1 dt lin    where dt is the sampling interval and lin the

number of header lines to be skipped. Will read any ascii data.

The last character of the last data line must be a <cr>, or the last

sample will be skipped.

lim  n: limit the length of the input series to n samples

end: end of input file, start the inversion for system parameters

Commands for signal processing:

ban  t0 m: gaussian FIR bandpass-filter with center period t0

and a total width of 2*m periods. The maximum amplitude of

the output signal is stored but the output signal itself

is discarded.

baf  same as ban, envelope of filtered signal is written into a file

avg  n: remove average (determined from first n samples)

dec  n: decimate to every n'th sample

dif  differentiate the signal

exp  n: remove exponential trend determined from first n samples

ext  determine extremal samples (signal is unchanged)

fac  f: multiply each sample by f

int  integrate the signal

lin  remove linear trend through first and last sample

nor  a: normalize amplitude to a

nop  a: normalize peak amplitude to a

nox  a: normalize extrema to +-a

nul  n1  n2: nulls the signal between samples n1 and n1 (inclusive)

pad  n; pad with zeroes to length n

pol  n: remove polynomial trend of degree n

rev  reverse signal in time (useful for backward filtering)

rms  determine rms amplitude (signal is unchanged)

skp  n: skip n samples

spl  clp,n1,n2,apex:  interpolate clipped samples (value >= clp) with

cubic spline from sample n1 to n2. clp=0 is replaced by a default

value of slightly less than 2**18. if apex is specified other than

zero, interplation is done with a parabola with specified height of

apex.

tap  n1  n2: signal is cosine- tapered between samples n1 and n2

(signal is unchanged between samples 1 and n1)

tre  n: remove linear trend (determined from first n samples)

twi  tmin1  tsec1  tmin2  tsec2: window defined by time after midnight

swi  sec1 sec2: window defined by seconds after first sample

win  n1  n2: time window from sample n1 to n2 (n1 becomes first sample)

Recursive (IIR) filters:

filters are specified by a three-character code for the type

and 1 to 4 parameters.

type: one of lp1, hp1, le1, he1, lp2, hp2, bp2, le2, he2, lpb, hpb

lp = low pass, hp= high pass, bp = band pass

parameters: corner period [,damping if second-order]

le = inverse low pass, he = inverse high pass (equalizers)

parameters: old corner period, [old damping], new corner

period, [new damping]

1 or 2: number of poles, i.e. order of denominator polynomial

b (in pace of 1 or 2):      Butterworth, specify period and order

Example: lpb 30. 6 is a 30 sec, sixth-order Butterworth low-pass filter.

c Additional type: sez, inverse second-order high-pass filter (inverse seismometer response)

##### Download

Download source code, executable, and test data from

http://www.software-for-seismometry.de/