| ## Name | ## DISPCAL |
| --- | --- |
| Author | Erhard Wielandt, Institute of Geophysics, University of Stuttgart, Richard-Wagner-Strasse 44, D - 70184 Stuttgart, E-mail: e.wielandt@t-online.de |
| Version | 2010 |

This program analyzes a single data file containing the output signal of a seismometer during repeated stepwise displacements. The basic concept is described in the ,Handbook' and ‘Manual' chapters on seismometry.

DISPCAL reads two input files: a parameter file dispcal.par and a data file whose name is specified there. It generates seven output files containing intermediate signals generated in the analysis. By plotting these, the user can follow the signal processing and determine reasonable values of the control parameters in the dispcal.par file. It also generates a text file dispcal.out with the screen output. An alternative name for the parameter file can be given as a runstring; the name of the .out file will correspond to the .par file.

The structure of the data file is as follows:

- one header line, arbitrary (will be echoed but not evaluated)

- additional header or comment lines whose first character is %

- one line containing the number of samples, the FORTRAN format in which they are listed, and the sampling interval. These three entries must be in the FORTRAN format (i10,a20,f10.x).

- data in the specified format

You may easily change this file structure in the source code if you want to read data in another format. DISPCAL can also read the ASL (Peterson) format.

The parameter file dispcal.par has the following structure:

'calibration of STS2 # 80444 on the Lennartz step table'

'data\disp.z'       name of data file

'STS2 #80444'  type and serial no. of seismo

120.           free period of seismometer (s)

0.707          fraction of critical damping

0.596          microvolts per count

0.903          displacement in millimeters per step

0.5            safety distance and percentage of quiet trace - see NOTE

0. 0        corner period (s) and order of low-pass filter (none if 0.,0)

.false.        print list of pulses

1.006          correction factor for STS2 with Q330HR, no analog Monitor

NOTE: First parameter: safety distance from the transitions in

seconds. This is also the minimum averaging length. The

entry 0 will cause the program to try five values from

0.125 to 1 s.

The first four numerical parameters and the last one must be known before DISPCAL is run.

The free period and damping of the seismometer can be obtained from an analysis with CALEX ore some other system-identification software. The microvolts per count refer to the calibration of the digital recorder and the millimeters of displacement to the stepwise-motion experiment. The correction factor compensates for the signal loss due to the source impedance of the sensor; it is normally slightly larger than unity. The other parameters help the program to decide which parts of the data record represent motion and which ones rest. If you are uncertain about the “safety distance”, enter zero and the program will try five value between 0.125 and 2 seconds. Select the one that produces the smallest rms scatter. The program expects to find at least 5 seconds without motion at the beginning of the input signal.

When the signal-to-noise ratio is poor, the program may not be able to separate the intervals where motion occurs from those without. You can help the program by inserting a second parameter ‘perc’ after the “safety distance”, in the same line and separated by a blank. It specifies the percentage of the total duration that should be considered as motion-free (typically about 80%). A plot of the signal dispcal.str (see below) will help you to make the right choice; positive values of dispcal.str indicate intervals of motion.

Accelerometers can also be calibrated on a step table with DISPCAL. The corner period in line 4 of the dispcal.par file must be set to zero in this case. Note that accelerometers respond more strongly than seismometers to high-frequency ground noise. It may therefore be necessary to activate the built-in low-pass filter routine in line 9 of the parameter file.

The 2017 version of DISPCAL can also deconvolve signals recorded with a first-order high-pass response to ground velocity (the normal response of a seismometer is second-order). When the damping is given as zero, the corner period is understood as that of a first-order high-pass filter.

No correction for the frequency response is made when the corner period is set to a negative value. This is useful when the response is unconventional so that DISPCAL cannot handle it. You must then remove the response from the data with another method before calling DISPCAL.

The line with the numerical damping must be present even if the free-period parameter is zero or negative, but the value is then ignored.

Be sure to use the correct source (RS) and recorder (RR) impedances when calculating the correction factor (RR+RS)/ RR.

Concept of the analysis:

Straight segments of the deconvolved velocity and displacement traces represent the rest phase while during the motion these traces are curved. The program will then deconvolve the trace and remove a higher-order trend from it as specified by the ,,degree" parameter. Then it will try to distinguish between intervals of motion and rest. For this purpose straight lines are fitted to suitable sections of the velocity trace and depending on the residual these sections are classified as curved or straight. After identifying the straight segments (i.e., time intervals without motion) the program discards part of each “straight” segment in order to keep a minimum distance from the transients. The result is stored in a logic (yes-no) signal which is written into the file dispcal.str. By inspecting a plot of this signal the user may check whether the straightness parameters were correctly chosen.

In a second step of trend removal only the straight segments of the deconvolved velocity are commonly detrended to the same order as before so that they can be used as a baseline. In a third step, straight lines are fitted to each segment of the baseline, and the trend in between is interpolated. When the trend has finally been piecewise removed, the velocity trace is integrated into displacement. The magnitude of the steps in the displacement trace is measured and by comparison with the known mechanical displacement, the generator constant of the seismometer is calculated for each step.

The program first averages the generator constant over all identified steps. Then it does error statistics and eliminates those steps that contribute most to the variance. The final result does therefore not depend on all steps being correctly identified, and is relatively insensitive to incorrectly chosen parameters.

The following intermediate signals are written into files:

*dispcal.bol    the logic signal: 0 = rest, 1 = motion

dispcal.dat     a copy of the original (broadband-velocity) data.

dispcal.dis     the restored displacement

dispcal.res     the residual velocity in the intervals of rest

*dispcal.srt     a sorted version of the straightness signal

dispcal.str      the straightness signal, from which periods of motion and rest are determined

dispcal.vel     The inversely filtered and detrended velocity trace

dispcal.stp      the scatter of the individual steps around the mean (repeatability)

*  presently disabled. Can be enabled in the source code, which must then be recompiled.

These signals can be plotted by typing WINPLOT. A plot-parameter file  ‘winplot.par’ is automatically generated, so the signals can be plotted by simply calling WINPLOT after DISPCAL. The parameter file may be edited if the plot does not fit to the screen or must otherwise be modified. It is strongly recommended to plot and check the signals for each experiment, to make sure that the motion of the table was properly restored from the recorded signal.

A protocol of the whole analysis is printed on the screen and written into the file dispcal.out.

Also, a separate list of the individual steps is produced as file ‘disptable.lst’.

Download source code, executable, and test data from

http://www.software-for-seismometry.de/

Another version of this program:

DISPCALX3 determines crosstalk in a three-component seismometer when only one com-ponent is deliberately excited on a step or shake table. By repeating the experiment for all three components the complete 3x3 matrix of crosstalk ratios is obtained.