### Exercise 5.5a: Interpreting Poles and Zeros from SEED headers

1  Aim

The complex transfer function (or the related complex frequency response) of the analog part of a seismograph is a rational function of frequency. Such functions can be specified by corner frequencies and damping constants, by polynomial coefficients, or by their poles and zeros. The latter method is chosen in the IRIS SEED data volumes. For each data channel of each station, the data header contains a list of poles and zeros of the transfer function together with some auxiliary information. IRIS supplies a software library 'evalresp' for extracting and interpreting these parameters. The exercise aims at making you familiar with interpreting poles and zeros in terms of  the amplitude response versus frequency.

2  Task

Interpret one or more of the attached SEED headers with respect to the analog part of the seismograph. Sketch the amplitude response for one of the stations as a Bode-diagram on double logarithmic paper. (The digital part is usually of minor interest since it is supposed to have a flat amplitude response and zero phase delay.) Does the header describe a very broadband, broadband or narrowband system? Note that the answer does not only depend on the mathematical form of the response but also on the definition of the input signal - displacement, velocity or acceleration. A broadband seismograph is supposed to have a broadband response to velocity but a broadband accelerometer has a broadband response to acceleration. Be careful with the units - some headers refer to Hertz rather than radians/sec. Check also whether the poles and zeros refer to the Laplace transform or Fourier transform. Can you guess which type of sensor is used? Are the constants nominal or were they determined from an individual calibration?

A little computer program POLZERO in BASIC will be made available to you to do the numerical conversions and to plot the amplitude response (more details are given in the program description). Use this program to analyze some more of the SEED headers. The stations are:

KIP (Kipapa, Hawaii)

KMI (Kunming, China)

XAN (Xi'an, China)

KIP

KMI

XAN

Solutions (use POLZERO for amplitude plots):

KIP   		velocity very broadband, lower corner 360 s, upper corner 0.2 s

Obviously an older STS1-VBB seismometer. No extra filters.

Nominal parameters.

KMI    	narrowband LP as a displacement sensor, but better characterized as

a long-period acceleration sensor. Response is flat to acceleration

from 30 s to 600 s. The sensor must be an old STS1 (20 s). A 6th-order

Butterworth low-pass filter limits the bandwidth at 30 s; this would

today be done with digital filters in the recorder. Nominal parameters.

XAN		velocity broadband, lower corner 120 s, upper corner 44 Hz.

Probably an STS2 or a CMG3-T seismometer. Additional low-pass

filter at 77 Hz. Parameters were probably measured.