Type; # S-13, GS-13. Manufacturer; # Geotech Instruments LLC. Address; 10755 Sanden Drive, Dallas, Texas 75238-1336, USA. Phone; (001) 214 221 0000. E-mail; info@geoinstr.com. Homepage; www.geoinstr.com/. Category; # Electromagnetic single-component, short-period seismometer, convertible between horizontal and vertical. Flat response; velocity 1 Hz (adjustable) to >100 Hz. Resolution: NLNM; 0.07 to 40 Hz with low-noise preamplifier (S-13)* 0.05 to 150 Hz with low-noise preamplifier (GS-13)*. within 10 dB of NLNM. within 20 dB of NLNM; above 0.04 Hz (GS-13)*. Operating range; large, limited by preamplifier. Generator constant; 630 Vs / m with 3600 Ohms coil (S-13) 2180 Vs / m with 3600 Ohms coil (GS-13). Output; differential, unlimited. Weight; 14 kg. Moving mass is 5 kg.. Size; 17 cm dia., 38 cm high. Power; passive. Manual control; mass lock, centering, conversion Z / H. Remote control; calibration. Remote diagnostic; -. Accessories; -. Typical installation; observatory, short-period. Data sheet; http://www.geoinstr.com/ds-s13.pdf. Remarks; A widely used high-performance, short-period seismometer. For very quiet sites, Model GS-13 is the best choice. It resolves high-frequency signals much better than any broadband-feedback seismometer. See publications by Riedesel et al. (BSSA 80,6) and by Rodgers (BSSA 83,2 and 84,1) for information on preamplifier design and noise.. * based on circuit analysis with op-amp LT1007 as a preamplifier. Low-noise model may be inaccurate above 10 Hz.