Type; # Le-3d lite. Manufacturer; # Lennartz Electronic. Address; Bismarckstrasse 136, D-72072 Tuebingen, Germany. Phone; ++49 7071 93550. E-mail; info@lennartz-electronic.de. Homepage; www.lennartz-electronic.de. Category; # Active short-period, three-component. Flat response; velocity 1 to 80 Hz. Resolution: NLNM; ---. within 10 dB of NLNM; 0.2 to 0.5 Hz (microseismic peak). within 20 dB of NLNM; 0.15 to 20 Hz. Operating range; ±13 mm/s below 4.5 Hz, decreasing at higher freq.. Generator constant; 400 Vs/m. Output; ±5 V single-ended. Weight; 1.8 kg. Size; 95 mm dia., 65 mm high. Power; 12 V, 0.1 W. Manual control; no mass lock required. Remote control; calibration. Remote diagnostic; -. Accessories; -. Typical installation; monitoring networks, field deployment. Data sheet; http://www.lennartz-electronic.de/MamboV4.5.2/index.php? option=com_content&task=view&id=72&Itemid=55. Remarks; A very small and rugged three-component, short-period, active seismometer. The sensor is a commercial 4.5 Hz geophone whose response is electronically extended by a negative damping resistance (a form of negative feedback). Not as sensitive as larger seismometers but good enough where ground noise is not extremely low. A single-component version Le-1d is also available. Higher performance is offered by the Le-3d/5s and Le-3d/20s..