Type; # Mark L4-3D (out of production). Manufacturer; # Mark Products (disappeared). Address. Phone. E-mail. Homepage. Category; # Passive electromagnetic three-component short-period. Flat response; velocity 1 Hz to >100 Hz (2 Hz version was available). Resolution: NLNM; 0.12 to 12 Hz with low-noise preamplifier*. within 10 dB of NLNM; 0.07 to 35 Hz with low-noise preamplifier*. within 20 dB of NLNM; 0.05 to 100 Hz with low-noise preamplifier*. Operating range; large, limited by preamplifier. Generator constant; 275 Vs/m with 5000 Ohms coil. output; differential, unlimited. Weight; 12 kg. Moving mass is 1 kg per component.. Size; 24 cm dia. , 24 cm high. Power; -. Manual control; no mass lock required. Short signal coils and tilt the instrument for transportation.. Remote control; -. Remote diagnostic; -. Accessories; -. Typical installation; temporary installations, field work. Data sheet; http://touro.ligo-la.caltech.edu/~ospjeld/HEPI_Documen tation_Website/HEPI_Main_Page/pdfs/L4.pdf. Remarks; A still widely used short-period seismometer for field work. See publications by Riedesel et al. (BSSA 80,6) and by Rodgers (BSSA 83,2 and 84,1) for information on preamplifier design and noise. The Mark L4-3D is out of production but OYO-Geospace offers a comparable three-component package, the “Seis-Monitor” based on the GS-1 geophone.. * based on circuit analysis with op-amp LT1007 as a preamplifier. Low-noise model may be inaccurate above 10 Hz.