Type; # STS-1 VBB (out of production, see remarks below). Manufacturer; # Streckeisen AG. Address; Dättlikoner Str. 5, CH-8422 Pfungen, Switzerland. Phone; ++41 52 315 2161. E-mail; none. Homepage; none. Category; # Force-balance VBB, separate Z and H components. Flat response; velocity 2.67 mHz (360 sec) to 10 Hz. Resolution: NLNM; originally <0.3 mHz to 3 Hz, later versions up to 8 Hz*. within 10 dB of NLNM; 0.1 mHz to 13 Hz*. within 20 dB of NLNM; <0.1 mHz to 20 Hz*. Operating range; ±8 mm/s. Generator constant; 2400 Vs/m. Output; differential, ±10 V per line. Weight; 4 kg (vert.); 5.5 kg (hor.). Size; 12 * 17 * 18 cm (vert.); 20 cm dia. and 16 cm high (hor.) Shields require a space of 50 * 50 cm, 60 cm high.. Power; ±15 volts, 3.5 watts per component. Manual control; mass lock (pins and screws inserted by hand). Remote control; mass centering, calibration. Remote diagnostic; mass position. Accessories; feedback electronics are separate. Magnetic shield (for vert. comp. only), aluminum shield, glass base plate, vacuum glass bell included. Monitor (breakout) box optional.. Typical installation; observatory. Data sheet; www.software-for-seismometry.de/textfiles/Seismometry/. Remarks; Operated under partial vacuum; no electronics inside the sensor. This seismometer made the very-broad-band velocity response popular, and is the only seismometer to resolve minimum ground noise throughout the long-period seismic band. (Some say it defines the low-noise model.) The STS1 is no longer manufactured by Streckeisen AG but Metrozet LLC (USA) have developed a modern, remotely controllable replacement for the feedback electronics and a modified mechanical sensor. The new system has the excellent low-frequency resolution of the original STS-1 but is easier to install and maintain.. * based on Ringler and Hutt 2010. Resolution above 10 Hz is extrapolated.