Type; # STS-2 (out of production; # replaced by STS 2.5). Manufacturer; # Streckeisen AG. Address; Dättlikoner Str. 5, CH-8422 Pfungen, Switzerland. Phone; ++41 52 315 2161. E-mail; none. Homepage; none. Category; # Force-balance VBB, symmetric-triaxial. Flat response; velocity 8.33 mHz (120 sec) to 50 Hz. Resolution: NLNM; 30 mHz to 8 Hz 1 3 mHz to 20 Hz 3. within 10 dB of NLNM; <1 mHz to 20 Hz 1 0.3 mHz to 35 Hz 3. within 20 dB of NLNM; <0.1 mHz to 50 Hz 1 <0.1 mHz to 50 Hz 3. Operating range; ±13 mm/s. Generator constant; 1500 Vs/m (high-gain version: 20 000 Vs/m). Output; differential, ±10 V per line. Weight; 13 kg. Size; 23 cm dia., 26 cm high. Power; 9 to 28 volts, 1.5 watts, low-power version was available. Manual control; mass lock. Remote control; centering, short-period mode, calibration, select UVW positions or UVW broadband signals. Remote diagnostic; mass positions or UVW signals (analog). Accessories; A “Host Box” with power conditioning and two parallel signal connectors included. Typical installation; observatory and vault. Needs additional protection in humid or dirty environments.. Data sheet; http://www.passcal.nmt.edu/content/instrumentation/sensors/broadband-sensors/sts-2-bb-sensor Manual: http://www.passcal.nmt.edu/webfm_send/488. Remarks; Tightly-fitting thermal isolation required for optimum performance. Until recently, the STS-2 was one of the two most widely used VBB instruments (the other being the CMG3). Three versions (generations) of the internal electronics exist. Production ended in 2010. The successor model STS 2.5 has the same housing but completely redesigned mechanical sensors and electronics, and can be remotely controlled with a laptop computer. Its performance is similar to that of the third-generation STS-2.. 1 first generation, 3 third-generation, high-gain STS-2, based on Ringler and Hutt 2010. Resolution above 10 Hz is extrapolated.