Type; # STS 2.5. Manufacturer; # Streckeisen GmbH; exclusively sold by Quanterra, USA. Address; St: Dättlikoner Str. 5, CH-8422 Pfungen, Switzerland Q: 325 Ayer Rd, Harvard, Ma. 01451, USA. Phone; ++41 52 315 6700 (Streckeisen) ++1 978 772 4774 (Quanterra). E-mail; bkemp@quanterra.com. Homepage; www.quanterra.com; www.kinemetrics.com/p-163-Home.aspx. Category; # Force-balance VBB, symmetric-triaxial. Flat response; velocity 8.33 mHz (120 sec) to 50 Hz. Resolution: NLNM; 3 mHz to 20 Hz 3. within 10 dB of NLNM; 0.3 mHz to 35 Hz 3. within 20 dB of NLNM; 0.1 mHz to 50 Hz 3. Operating range; ±13 mm/s. Generator constant; 1500 Vs/m. Output; differential, ±10 volts per line. Weight; 12 kg. Size; 23 cm dia., 26 cm high. Power; 9 to 28 volts, 0.5 watts. Manual control; -. Remote control; mass lock, centering, calibration, short-period mode. Remote diagnostic; mass position, state-of health parameters including temperature, internal humidity, tilt, serial number, supply voltage. Accessories; a “Host Box” with power conditioning and LED status indicators is included.. Typical installation; Observatory, vault. Needs additional protection in humid or dirty environments.. Data Sheet; http://www.kinemetrics.com/uploads/PDFs/STS-2_5.pdf. Remarks; Tightly-fitting thermal isolation required for optimum performance. The STS 2.5 has the same housing as the STS-2 but completely redesigned mechanical sensors and electronics. Remote control with a laptop computer. The STS 2.5 has not yet been extensively tested but appears to perform equally or slightly better than a third-generation STS-2.. 3 self-noise data are from a third-generation, high-gain STS2, based on Ringler and Hutt 2010. Resolution above 10 Hz is extrapolated.